## AIM
Test# | Desc | Equipment
9000 | # HPIB Sales Rack Verification
9006 | # Null Multi-module test
9002 | # Analog Fixture Present
9255 | # DC SRC voltage accuracy and range | DMM, ANA FIX
9256 | # DC SRC voltagey headroom test | DMM, ANA FIX
9260 | # DC SRC output resistance | DMM, ANA FIX
9265 | # DC SRC positive overvoltage limit accuracy | DMM, FG, ANA FIX | FG provide DC start at 10V(incremental 1.024V), to cause OV on Asru DC SRC(around 13V). Inremental reduce to half when OV detected & DMM log the actual value
9266 | # DC SRC negative overvoltage limit accuracy | DMM, FG, ANA FIX | FG provide DC start at -10V(incremental 1.024V), to cause OV on Asru DC SRC(around -13V). Inremental reduce to half when OV detected & DMM log the actual value
9267 | # DC SRC positive overvoltage protection | DMM50, FG, ANA FIX | FG provide DC start at 10V, then set to 20V to cause OV on Asru DC SRC. Inremental reduce to half when OV detected & DMM log the actual value
9268 | # DC SRC negative overvoltage protection | DMM50, FG, ANA FIX | FG provide DC start at -10V, then set to -20V to cause OV on Asru DC SRC. When OV detected & DMM log the actual value
9270 | # DC SRC tri-state spike test | Counter, ANA FIX | Counter is setup to 50OHM & DC couple to sample the peak volatge of Asru DC source at tri-state for +5V & -5V
9275 | # DC SOURCE OUTPUT CURRENT TEST | PS, ANA FIX
9274 | # DC SOURCE OUTPUT CURRENT TEST FOR NASRU | PS, ANA FIX
9277 | # DC SRC max current (HIGH LEVEL ONLY) | PS, ANA FIX
9305 | # AUX SRC voltage accuracy and range | DMM, ANA FIX
9306 | # AUX SRC voltage headroom test | DMM, ANA FIX
9310 | # AUX SRC output resistance | DMM, ANA FIX
9315 | # AUX SRC positive overvoltage limit accuracy | DMM, FG, ANA FIX | FG provide DC start at +/-10V(incremental 1.024V), to cause OV on Asru DC SRC(around 13V). Inremental reduce to half when OV detected & DMM log the actual value
9316 | # AUX SRC negative overvoltage limit accuracy | DMM, FG, ANA FIX | FG provide DC start at +/-10V(incremental 1.024V), to cause OV on Asru DC SRC(around -13V). Inremental reduce to half when OV detected & DMM log the actual value
9317 | # AUX SRC positive overvoltage protection | DMM50, FG, ANA FIX | FG provide DC start at 10V, then set to 20V to cause OV on Asru DC SRC. Inremental reduce to half when OV detected & DMM log the actual value
9318 | # AUX SRC negative overvoltage protection | DMM50, FG, ANA FIX | FG provide DC start at -10V, then set to -20V to cause OV on Asru DC SRC. When OV detected & DMM log the actual value
9320 | # AUX SRC tri-state spike test | Counter, ANA FIX | Counter is setup to 50OHM & DC couple to sample the peak volatge of Asru AUX source at tri-state for +5V & -5V
9325 | # AUX SOURCE OUTPUT CURRENT TEST | PS, ANA FIX
9324 | # AUX SOURCE OUTPUT CURRENT TEST FOR NASRU | PS, ANA FIX
9327 | # AUX SRC max current (HIGH LEVEL ONLY | PS, ANA FIX
9355 | # AC SRC voltage accuracy and range | DMM, ANA FIX
9360 | # AC SRC dc offset voltage accuracy | DMM, ANA FIX
9367 | # AC source frequency accuracy and jitter | DMM, ANA FIX
9370 | # AC SRC output resistance | DMM, ANA FIX
9371 | # AC SRC THD | DMM, ANA FIX
9372 | # AC SRC THD FOR NASRU | DMM, ANA FIX
9380 | # ARB SRC THD Dual Frequency Mode | DMM, ANA FIX
9383 | # ARB THD | DMM, ANA FIX
9386 | # ARB SRC square wave accuracy, % overshoot, slew rate | DMM, DMM50, ANA FIX
9387 | # ARB SRC triangle wave rise/fall time and linearity | DMM, DMM50, ANA FIX
9391 | # ARB SOURCE SINE WAVE FREQUENCY ACCURACY USING DMM | DMM, ANA FIX
9392 | # ARB SOURCE SINE WAVE PERIOD JITTER MEASUREMENT USING DMM | DMM, ANA FIX
9393 | # ARB SOURCE SQUARE WAVE PERIOD JITTER MEASUREMENT USING COUNTER | Counter, ANA FIX | Counter is setup to measure ASRU ARB source square wave frequency at 10Hz, 100Hz, 1KHz, 10KHz & 20KHz for jitter measurement
9394 | # ARB SOURCE SQUARE WAVE FREQUENCY ACCURACY USING COUNTER | Counter, ANA FIX | Counter is setup to measure ASRU ARB source square wave frequency at 10Hz, 100Hz, 1KHz, 10KHz & 20KHz for accuracy
9410 | # MOA output voltage range | DMM, PS, ANA FIX
9415 | # MOA positive overvoltage limit accuracy | DMM, FG, ANA FIX | FG provide DC start at +/-10V(incremental 1.024V), to cause OV on Asru DC SRC(around 13V). Inremental reduce to half when OV detected & DMM log the actual value
9416 | # MOA negative overvoltage limit accuracy | DMM, FG, ANA FIX | FG provide DC start at +/-10V(incremental 1.024V), to cause OV on Asru DC SRC.(around -13V) Inremental reduce to half when OV detected & DMM log the actual value
9420 | # MOA dc output resistance | DMM, ANA FIX
9427 | # MOA maximum output current (HIGH LEVEL ONLY) | PS, ANA FIX
9440 | # MOA CMRR | DMM, ANA FIX
9455 | # DC DETECTOR accuracy range 8 | DMM, ANA FIX
9456 | # DC DETECTOR accuracy range 7 | DMM, ANA FIX
9457 | # DC DETECTOR accuracy range 6 | DMM, ANA FIX
9458 | # DC DETECTOR accuracy range 5 | DMM, ANA FIX
9459 | # DC DETECTOR accuracy range 4 | DMM, ANA FIX
9460 | # DC DETECTOR accuracy range 3 | DMM, ANA FIX
9461 | # DC DETECTOR accuracy range 2 | DMM, ANA FIX
9462 | # DC DETECTOR accuracy range 1 | DMM, ANA FIX
9463 | # DC DETECTOR accuracy range 0 | DMM, ANA FIX
9451 | # DC DETECTOR ATTENUATOR ACCURACY (range 8) | DMM, PS, ANA FIX
9452 | # DC DETECTOR ATTENUATOR ACCURACY (range 7) | DMM, PS, ANA FIX
9453 | # DC DETECTOR ATTENUATOR ACCURACY (range 6) | DMM, PS, ANA FIX
9441 | # DC DETECTOR accuracy range 8-NASRU | DMM, ANA FIX
9442 | # DC DETECTOR accuracy range 7-NASRU | DMM, ANA FIX
9443 | # DC DETECTOR accuracy range 6-NASRU | DMM, ANA FIX
9444 | # DC DETECTOR accuracy range 5-NASRU | DMM, ANA FIX
9445 | # DC DETECTOR accuracy range 4-NASRU | DMM, ANA FIX
9446 | # DC DETECTOR accuracy range 3-NASRU | DMM, ANA FIX
9447 | # DC DETECTOR accuracy range 2-NASRU | DMM, ANA FIX
9448 | # DC DETECTOR accuracy range 1-NASRU | DMM, ANA FIX
9449 | # DC DETECTOR accuracy range 0-NASRU | DMM, ANA FIX
9465 | # DETECTOR true rms accuracy over frequency at 7.071 Vrms | DMM, FG, ANA FIX | FG provide AC signal(SINE) 7.07Vrms or +/-10Vpeak with freq Fast mode: 50, 200, 400, 5000, 10000, and 19000 Hz Normal mode: 10, 50, 300, 3000, 10000, and 19000 Hz FG supplied volatge is read back by DMM
9466 | # DETECTOR true rms accuracy over voltage ranges at 19 KHz | DMM, ANA FIX
9470 | # DC DETECTOR input resistance | DMM, ANA FIX
9471 | # DETECTOR CMRR | Not require
9472 | # DETECTOR CMRR (Attenuated) | Not require
9475 | # DC DETECTOR Quadrature Rejection - Loaded Source | Not require
9474 | # DC DETECTOR Quadrature Rejection - Loaded Source FOR NASRU | Not require
9476 | # DC DETECTOR Source Crossover Quadrature Rejection | Not require
9477 | # DC DETECTOR Source Crossover Quadrature Rejection FOR NASRU | Not require
9480 | # DC DIGITIZER VOLTAGE ACCURACY | DMM, ANA FIX
9481 | # DC DIGITIZER NOISE OVER RANGES | Not require
9482 | # DC DIGITIZER NOISE OVER RANGES | Not require

## ANA
Test# | Desc | Equipment
9006 | # Null Multi-module test
9002 | # Analog Fixture Present
9000 | # GPIB Sales Rack verification
9500 | # Borg command module test
9010 | # Sales Rack Offset Capacitance Test
9011 | # Load test for the channel #7 UTILITY CARD
9012 | # Load test for the channel #8 UTILITY CARD
9200 | # [PINCARD] Series Xn Bus Resistance | DMM, ANA FIX
9201 | # [PINCARD] Series XG Bus Resistance | DMM, ANA FIX
9207 | # [PINCARD] PARALLEL Xn BUS RESISTANCE USING MOA | Not require
9208 | # [PINCARD] PARALLEL Xn BUS RESISTANCE USING MOA ASRU & MB ONLY | Not require
9209 | # [PINCARD] PARALLEL Xn TO XG BUS RESISTANCE USING MOA | Not require
9210 | # [PINCARD] Series Xn Bus Inductance | LCR, ANA FIX
9211 | # [PINCARD] Series Xn Bus Inductance | LCR, ANA FIX
9217 | # [PINCARD] MOA Xn to GROUND RESISTANCE | Not require
9220 | # [PINCARD] Parallel Xn Bus Capacitance | LCR, ANA FIX
9221 | # [PINCARD] Parallel XG Bus Capacitance | LCR, ANA FIX
9230 | # [PINCARD] Xn Bus Leakage Current | DMM, DIG FIX
9900 | # [TWINPIN] Series Xn Bus Resistance | DMM, ANA FIX
9901 | # [TWINPIN] Series XG Bus Resistance | DMM, ANA FIX
9907 | # [TWINPIN] PARALLEL Xn BUS RESISTANCE USING MOA | Not require
9909 | # [TWINPIN] PARALLEL Xn TO XG BUS RESISTANCE USING MOA | Not require
9910 | # [TWINPIN] Series Xn Bus Inductance | LCR, ANA FIX
9911 | # [TWINPIN] Series Xn Bus Inductance | LCR, ANA FIX
9917 | # [TWINPIN] MOA Xn to GROUND RESISTANCE | Not require
9920 | # [TWINPIN] Parallel Xn Bus Capacitance | LCR, ANA FIX
9921 | # [TWINPIN] Parallel XG Bus Capacitance | LCR, ANA FIX
9930 | # [TWINPIN] Xn Bus Leakage Current | DMM, DIG FIX

## DIG
Test# | Desc | Equipment
9005 | # Read the delay constants for this sales rack
9006 | # Run a multi-module test to assure everything runs in serial
9000 | # HPIB Sales Rack Verification
9003 | # Check for the digital fixture
9020 | # Calculate the ASRU 5ohm load and dmm/osc ratio for each ch | DMM, DIG FIX
9021 | # Calculate the ASRU 24ohm load | DMM, DIG FIX
9050 | # TIC verify - Frequency Mode | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp(20MHz & 60MHz) to sync PG. POS & NEG pulse of PG output is measured by counter. The difference of TIC & counter measurements are the log result
9052 | # TIC verify - TI MODE | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp(20MHz & 60MHz) to sync PG. POS & NEG pulse of PG output is measured by counter. The difference of TIC & counter measurements are the log result
9053 | # TIC verify - TI MODE | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp20MHz & 60MHz to sync PG. POS & NEG pulse of PG output is measured by counter. The difference of TIC & counter measurements are the log result
9055 | # USER CLOCK GENERATION RANGE, ACCURACY, & JITTER | Counter, DIG FIX
9056 | # USER CLOCK GENERATION RANGE, ACCURACY, & JITTER (50MHz) | Counter, DIG FIX
9060 | # SYNC TO CLOCK RANGE AND JITTER | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp(5MHz, 10MHz, 20MHz & 40MHz) to sync PG. PG clock output to CLKRB, Freq & jitter of TCLK & PG is measured by counter.
9061 | # SYNC TO CLOCK RANGE AND JITTER (60Mz test) | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp(5MHz, 10MHz, 20MHz & 40MHz) to sync PG. PG clock output to CLKRB, Freq & jitter of TCLK & PG is measured by counter.
9065 | # SEQUENCER MARGIN TEST | FG, PG, DIG FIX | FG provide sinewave at 1.0Vpp(5MHz) to sync PG. PG output clock to CLKRB to trigger TCLK MUX & sequencer circuit
9071 | # USER TIME AT 12.5MHZ GUARANTEED > 30nS | DIG FIX
9075 | # EVENT TRIGGER VERIFICATION | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp(5MHz) to sync PG. PG output clock to CLKRB to sync ET circuit
9080 | # Boundry scan Tap Port (XTP)
9081 | # Firmware revision of sequencer (XTP)
9082 | # Firmware revision of sequencer (XTP)
9090 | # TIC VERIFY THROUGH PIN CARD - FREQUENCY MODE | FG, PG, DIG FIX | FG provide sinewave at 1.0Vpp(20MHz) to sync PG
9092 | # TIC VERIFY THROUGH PIN CARD - TI MODE | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp(20MHz) to sync PG
9095 | # USER CLOCK GENERATION RANGE, ACCURACY, & JITTER THROUGH PIN | Counter, DIG FIX
9100 | # DRIVER OUTPUT VOLTAGE RANGE AND ACCURACY (5v DRIVE HIGH) | DMM, DIG FIX
9101 | # [PINCARD] DR OUT V RANGE AND ACCURACY (-3.5v DRIVE HIGH) | DMM, DIG FIX
9102 | # DRIVER OUTPUT VOLTAGE RANGE AND ACCURACY (-3.5v DRIVE LOW) | DMM, DIG FIX
9103 | # [PINCARD] DR OUT V RANGE AND ACCURACY (5v DRIVE LOW) | DMM, DIG FIX
9104 | # GROUND DIFFERENCE TEST (PIN CARD TO AIM SWD GND) | DMM, DIG FIX
9106 | # [PINCARD] DRIVER OUTPUT SLEW RATE (100V/uS) | Counter, DIG FIX
9107 | # [PINCARD] DRIVER OUTPUT SLEW RATE (50V/uS) | Counter, DIG FIX
9111 | # DRIVER OUTPUT PULSE WIDTH (0.2 to 3.5v measured at 1.5v) | OSC, DIG FIX
9125 | # INTERMODULAR DRIVER EDGE PLACEMENT ACCURACY AND SKEW | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp(5MHz) to sync PG
9132 | # [PINCARD] DRIVER TRI-STATE LEAKAGE CURRENT | DMM, DIG FIX
9133 | # DRIVER LIMIT POWER PROTECTION - SOURCING (+4v, 4.5 ohms) | OSC, DIG FIX
9134 | # DRIVER LIMIT POWER PROTECTION - SINKING (-3v, 4.5 ohms) | OSC, DIG FIX
9135 | # [PINCARD] DRIVER OUTPUT IMPEDANCE (SOURCING 700mA) | DMM50, DIG FIX
9136 | # [PINCARD] DRIVER OUTPUT IMPEDANCE (SINKING 500mA) | DMM50, DIG FIX
9140 | # MAXIMUM DRIVE CURRENT (>700mA Sourcing) | DMM50, DIG FIX
9141 | # MAXIMUM DRIVE CURRENT (>500mA Sinking) | DMM50, DIG FIX
9142 | # LOW LEVEL CURRENT LIMIT SOURCING (60mA) | DMM, DIG FIX
9143 | # LOW LEVEL CURRENT LIMIT SINKING (40mA) | DMM, DIG FIX
9145 | # DRIVER LIMIT POWER PROTECTION - SOURCING ( 3.5v 24.5ohms) | OSC, DIG FIX
9146 | # DRIVER LIMIT POWER PROTECTION - SINKING (-2.5v 24.5ohms) | OSC, DIG FIX
9147 | # [PINCARD] DR POSITIVE OV LIMIT ACCURACY | DMM, DIG FIX
9148 | # [PINCARD] DR NEGATIVE OV LIMIT ACCURACY | DMM, DIG FIX
9150 | # RECEIVE VOLTAGE ACCURACY (+5v) | DMM, DIG FIX
9151 | # RECEIVE VOLTAGE ACCURACY (+4v) PULLED_UP | DMM, DIG FIX
9152 | # RECEIVE VOLTAGE ACCURACY (-3.5v) | DMM, DIG FIX
9153 | # RECEIVE VOLTAGE ACCURACY (-2.5v) PULLED_DOWN | DMM, DIG FIX
9160 | # RECEIVER INPUT CURRENT SOURCING | DMM, DIG FIX
9161 | # RECEIVER INPUT CURRENT SINKING | DMM, DIG FIX
9162 | # RECEIVER INPUT LEAKAGE CURRENT | DMM, DIG FIX
9165 | # RECEIVER INPUT CLAMP PULL UP | DMM, DIG FIX
9166 | # RECEIVER INPUT CLAMP PULL DOWN | DMM, DIG FIX
9167 | # RECEIVER INPUT CLAMP @ 0 VOLTS | DMM, DIG FIX
9182 | # [PINCARD] RX POSITIVE OV LIMIT ACCURACY | DMM, DIG FIX
9183 | # [PINCARD] RX NEGATIVE OV LIMIT ACCURACY | DMM, DIG FIX
9195 | # PROBE DC VOLTAGE RANGE & ACCURACY AT 5V | DMM, DIG FIX
9196 | # PROBE DC VOLTAGE RANGE & ACCURACY AT -3.5V | DMM, DIG FIX
9800 | # DRIVER OUTPUT VOLTAGE RANGE AND ACCURACY (5v DRIVE HIGH) | DMM, DIG FIX
9801 | # [TWINPIN] DR OUT VOLTAGE RANGE AND ACCURACY (-3.0v DRIVE HIGH) | DMM, DIG FIX
9802 | # DRIVER OUTPUT VOLTAGE RANGE AND ACCURACY (-3.5v DRIVE LOW) | DMM, DIG FIX
9803 | # [TWINPIN] DR OUT VOLTAGE RANGE AND ACCURACY (4.5v DRIVE LOW) | DMM, DIG FIX
9804 | # GROUND DIFFERENCE TEST (PIN CARD TO AIM SWD GND) | DMM, DIG FIX
9811 | # DRIVER OUTPUT PULSE WIDTH (0.2 to 3.5v measured at 1.5v) | OSC, DIG FIX
9820 | # DRIVER EDGE PLACEMENT ACCURACY AND SKEW (0 & 100nS DELAYS) | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp(5MHz) to sync PG
9829 | # TRI-STATE PULLED_UP DRIVE LOW EDGE PLACEMENT ACCURACY AND SKEW | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp(5MHz) to sync PG
9830 | # TRI-STATE HIGH EDGE PLACEMENT ACCURACY AND SKEW | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp(5MHz) to sync PG
9832 | # [TWINPIN] DRIVER TRI-STATE LEAKAGE CURRENT | DMM, DIG FIX
9835 | # [TWINPIN] DR OUTPUT IMPEDANCE (SOURCING 700mA) | DMM50, DIG FIX
9836 | # [TWINPIN] DR OUTPUT IMPEDANCE (SINKING 500mA) | DMM50, DIG FIX
9840 | # MAXIMUM DRIVE CURRENT (>700mA Sourcing) | DMM50, DIG FIX
9841 | # MAXIMUM DRIVE CURRENT (>500mA Sinking) | DMM50, DIG FIX
9847 | # [TWINPIN] DR POSITIVE OV LIMIT ACCURACY | DMM, DIG FIX
9848 | # [TWINPIN] DR NEGATIVE OV LIMIT ACCURACY | DMM, DIG FIX
9850 | # RECEIVE VOLTAGE ACCURACY (+5v) | DMM, DIG FIX
9851 | # RECEIVE VOLTAGE ACCURACY (+4v) PULLED_UP | DMM, DIG FIX
9852 | # RECEIVE VOLTAGE ACCURACY (-3.5v) | DMM, DIG FIX
9853 | # RECEIVE VOLTAGE ACCURACY (-2.5v) PULLED_DOWN | DMM, DIG FIX
9860 | # RECEIVER INPUT CURRENT SOURCING | DMM, DIG FIX
9861 | # RECEIVER INPUT CURRENT SINKING | DMM, DIG FIX
9862 | # RECEIVER INPUT LEAKAGE CURRENT | DMM, DIG FIX
9871 | # [TWINPIN] RX EDGE PLACEMENT ACCURACY AND SKEW | FG, PG, Counter, DIG FIX | FG provide sinewave at 1.0Vpp(5MHz) to sync PG
9882 | # [TWINPIN] RX POSITIVE OV LIMIT ACCURACY | DMM, FG, DIG FIX | FG provide DC start at 13V, then set to 17.60V to cause OV on Asru DC SRC.
9883 | # [TWINPIN] RX NEGATIVE OV LIMIT ACCURACY | DMM, FG, DIG FIX | FG provide DC start at -13V, then set to -17.60V to cause OV on Asru DC SRC.